## Slide 1: 7. Measurements of lightning electric and magnetic fields

## Slide 2
Course Sections | Number of Lectures | HW Due (Canvas) | Test (Honor-lock) | Section in Course Packet | Chapter/Sections in Text (FOL)
5. Calculation of lightning electromagnetic fields | 4 | HW6 04/14 | T6 04/16 | 5 | 5, App. 3
6. Modeling of lightning processes | 3 |  |  | 6 | 6
7. Measurements of lightning electric and magnetic fields | 2 |  |  | 7 | 7
8. Lightning locating systems and lightning protection | 2 | - | - | 8,9 | 8,9
Review | 1 |  |  | 1-9 | 1-9, App. 1-4
Final Exam (via Honorlock) | - | - | Apr. 28 (W) 12:30-2:30 p.m. | 1-9 | 1-9, App. 1-4

## Slide 3
Illustration of the principle of operation of the flat-plate antenna. (a) Antenna without external circuit. (b) Antenna with external integrating capacitor C>>Ca. Drawing by Potao Sun.

## Slide 4
Norton equivalent circuit of electric field antenna shown along with the integrating capacitance and the input impedance of recorder (usually C>>Ca and C>>Cin). Drawing by Potao Sun.
Vout = (ε0AE/C) if Rin >> 1/(ωC

## Slide 5: For a unit-step-function E-field waveform, Vout = (ε0AE/C) exp(-t/(RinC)),which becomes Vout = (ε0AE/C) if
in the time domain: t << τ or t << RinC (practically impossible for quasi-static fields)
in the frequency domain: 1/(ωC) << Rin =>
f >> 1/(2πRinC) [lower frequency response]

## Slide 6
τ = 4 s, low gain
τ = 70 µs,
high gain
↓
↓

## Slide 7
Thevenin equivalent circuit of electric field antenna shown along with the integrating capacitance and the input impedance of recorder (usually C>>Ca and C>>Cin).
VT = IN (1/jωCa), where IN = ε0A jωE
For a unit-step-function E-field waveform,
Vout = (ε0AE/C) exp(-t/(RinC)),which becomes Vout = (ε0AE/C) if
- in the time domain: t << τ or t << RinC (practically impossible for quasi-static fields)
- in the frequency domain: 1/(ωC) << Rin => f >> 1/(2πRinC) [lower frequency response]

## Slide 8
Flat-plate antenna for measuring electric field (vertical cross-section and plan views) suitable for flush-mounting in the earth. Drawing by Potao Sun.

## Slide 9: Flat-plate antenna

## Slide 10
Elevated flat-plate antenna

## Slide 11: (a) Flat-plate antenna flush with ground (no field enhancement), (b) elevated flat-plate antenna, and (c) elevated flat-plate antenna on the top of building
For (a): Vout = (Gelectronics 𝐴𝜀0 /C ) E = E/Kflush and hence E = Kflush Vout,
where Kflush is the calibration factor. For (b) and (c): E = K Vout
Kflush is found theoretically, Gelevation experimentally, and Gbuilding numerically.
E = Kflush Vout
Vout = (Gelectronics Aε0/C) E

## Slide 12
Course Sections | Number of Lectures | HW Due (Canvas) | Test (Honor-lock) | Section in Course Packet | Chapter/Sections in Text (FOL)
5. Calculation of lightning electromagnetic fields | 4 | HW6 04/14 | T6 04/16 | 5 | 5, App. 3
6. Modeling of lightning processes | 3 |  |  | 6 | 6
7. Measurements of lightning electric and magnetic fields | 2 |  |  | 7 | 7
8. Lightning locating systems and lightning protection | 2 | - | - | 8,9 | 8,9
Review | 1 |  |  | 1-9 | 1-9, App. 1-4
Final Exam (via Honorlock) | - | - | Apr. 28 (W) 12:30-2:30 p.m. | 1-9 | 1-9, App. 1-4

## Slide 13
Signal to be
measured:
E = 1 kV/m
Δt = 0.5 ms
Recorder:
Vout = 2.5 V (±5 V)
Rin = 1 MΩ
Cin = 30 pF
Design
parameters:
C and A
C = (ε0AE/Vout)
Δt << RinC

## Slide 14
14
Typical measured vertical electric field intensity (left column) and azimuthal magnetic flux density (right column) waveforms for first (solid line) and subsequent (dashed line) return strokes at distances of 1, 2, 5, 10, 15, 50, and 200 km. Adapted from Lin et al. (1979).

## Slide 15
(a) Illustration of the principle of operation of the electric field mill. Adapted from Bazelyan and Raizer (2000). (b) Electric field mill operating at Camp Blanding, Florida. Photograph by Joseph Howard.

## Slide 16
16
Vertical electric field, E(t), measurements inside or in the immediate vicinity of the lightning channel and current, I(t), measurements at the channel base can be used to estimate power per unit length, P(t) = E(t) I(t), as a function of time. Integrating this power over time, we can estimate lightning input energy per unit length, W(t).
E(t)
I(t)
Lightning channel
r
Lightning input energy is needed in a number of areas including the estimation of NOx produced by lightning, evaluation of lightning shock-wave effects, and testing of thunder generation theories.
Estimation of Input Energy in Rocket-Triggered Lightning Strokes
R(t) = E(t)/I(t)
r(t) = [σ π R(t)]-1/2
Resistance per unit length, R(t), and channel radius, r(t) can be also estimated.

## Slide 17
Illustration of Pockels effect. Adapted from Miki et al. (2002).

## Slide 18
Schematic representation of experimental setup used to measure two components of electric field 01-1.6 m from the lightning channel with Pockels sensors. Adapted from Miki et al. (2002).

## Slide 19
19
Electric Fields in the Immediate Vicinity of the Lightning Channel
Vertical and horizontal (radial) components of electric field in the immediate vicinity (within 0.1 to 1.6 m) of the triggered-lightning channel were measured with Pockels sensors at Camp Blanding, Florida, in 2000 (Miki et al., 2002)
2-meter strike rod
Rocket launcher
Pockels sensor:
KDP (KH2PO4) crystal 20 kV/m – 5 MV/m
50 Hz – 1 MHz

## Slide 20
20
Estimation of Input Energy in Rocket-Triggered Lightning Strokes
Flash 0013, stroke 1. a) Vertical electric field 10 cm from the lightning channel attachment point, b) current at the channel base, c) power per unit channel length, d) energy per unit channel length. Adapted from Jayakumar et al. (2006).
a) E-Field
b) Current
c) Power
d) Energy
0.5 GW/m
(peak)
3 kJ/m
(at 46 μs)
R(t) = 5 Ω/m, r(t) = 3 mm at power peak (σ = 104 S/m)

## Slide 21
Norton equivalent circuit of electric field antenna shown along with the integrating capacitance and the input impedance of recorder (usually C>>Ca and C>>Cin). Drawing by Potao Sun.
Vout = (ε0AE/C) if Rin >> 1/(ωC

## Slide 22
Illustration of the principle of operation of the loop antenna. Drawing by Potao Sun.

## Slide 23
Thevenin equivalent circuit of magnetic field antenna shown along with the integrating circuit and the input impedance of recorder (C>>Cin).
Vout = ABn/(RC), if the following 3 conditions are met:
R >> ωL (ω << R/L, upper frequency response),
Rin >> R (C discharges through R, not Rin), and
3) R >> 1/(ωC) (ω >> 1/(RC), lower frequency response).
(B = Bn)

## Slide 24
Norton equivalent circuit of magnetic field antenna shown along with the integrating circuit and the input impedance of recorder. Drawing by Potao Sun.

## Slide 25
Loop antenna with single-ended output developed by George Schnetzer. (a) Schematic diagram; (b) Thevenin equivalent circuit. Drawing by Potao Sun.